- Provide technical specifications of the laser or laser system and a brief description of the work to be performed with the laser (include a copy of the vendor's specification and classification, if available).

- Describe the facility/environment in which the laser or laser system will be used (research laboratory, teaching laboratory, medical/surgical laboratory, office, etc.)  Draw a map of the lab and include placement of LASER units.

- Describe the level of laser safety knowledge and training of the personnel working with the laser. Create records of training including names and dates trained. Also address the presence of any other personnel who may not work directly with the laser, but may be exposed to hazards in the laser work area.

- Describe the safety and control measures present currently in the facility.

- Describe the safety and control measures that will be added/implemented (dates) along with the laser installation and how those measures will be achieved, including any protective housing, warning signs, interlocks etc. 
( see engineering  and administrative controls below)

- Describe any special ancillary hazards such as toxic materials/fumes, electrical exposures, or compressed gases and specific control measures that will be implemented to control said hazards. (see handout)

- Provide a standard operating procedure (SOP) for general operation, maintenance, and service procedures.

- Will operation of this laser or laser system involve the presence or any exposure to the general public at any time (such as special tours) or any other unusual circumstances?

[ ] Yes , [ ] No ; if yes,  describe.

- Will operation of this laser or laser system involve using lasers for health care, medical, or surgical applications to animals or human patients? 
[ ] Yes , [ ] No ; if yes,  indicate what institutional reviews are required/obtained, and describe the procedure(s) in which the laser will be used.

Engineering Control Measures

Area Posting – Area is posted with required Hazard Warnings    		Y	N

Controlled access – Only Authorized personnel have key access to restricted area

Y	N

Interlocks on all areas access doors are present and functioning where required

Y	N

Laser Hazard Labels are posted on all equipment				Y	N

Non-laser Hazard Labels i.e. Chemical, electrical etc. are posted  where required

Y	N

Interlocks on Entrance doors 							Y	N

Protective Housing are present and functioning where required		Y	N

Service Access Panel (electrical) limits entry only to authorized personnel

Y	N

Key Switch Master de-energizes the beam					Y	N

Remote Interlock Connection where required is functional			Y	N

Beam Stop or Attenuator is in place and fully operational			Y	N

Activation Warning Systems are functional and provide clear indication of hazard

Y	N

Where any of the above answers are negative, describe the detailed steps taken to meet the regulatory requirement.

Administrative and Procedural Controls

Standard Operating Procedures

- Describe in detail which eye protection is to be worn- include specific wording as to which eyewear is to be worn and when. Be specific: i.e. the pair in the left drawer of the computer bench. This pair has brown-framed goggles with the orange colored lenses with the number 5 on the frame.  If more than one type is available provide details for which pair is for which laser.

- Describe in detailed sequentially numbered steps, the safe energizing of the unit. Include steps to be taken prior to energizing the lasing system – ensuring coolant system is operational, other systems integration, computers, etc   For easy reference, post a copy of these protocols at or near the instrument.

- Where applicable, describe in detailed sequentially numbered steps, the appropriate protocol for aligning the beam paying particular attention to beam stops and specular surfaces.

- Describe in detailed sequentially numbered steps, the safe operation of the unit.

- Describe in detailed sequentially numbered steps, the protocol to follow in the event of emergency i.e. water spill, accidental personal exposure, fire etc.

- Describe in detail, the information that new users will receive prior to them being authorized to use the unit, record the names of all individuals authorized to use the unit, record the training dates of all authorized users and how they were trained. i.e. UBC Laser Safety course September 11, 2003 plus User orientation by lab supervisor Dr. Fred Flintstone October 23, 2003

- Describe in detailed numbered steps the procedures for engaging the services of UBC Land and Buildings personnel in the laboratory. (This may be as simple as stating that under no circumstance is the laser to be operated during any cleaning or maintenance procedures)

- If the electrical system is to be manipulated by research personnel, describe the level of training, date trained.  Ensure capacitor discharge is included in the hazard awareness training.

- Will operation of this laser or laser system involve the presence or any exposure to the general public at any time (such as special tours) or any other unusual circumstances?
[ ] Yes , [ ] No ; if yes, contact the UBC Radiation Safety office on each  occasion prior to permitting access.

- Will operation of this laser or laser system involve using lasers for health care, medical, or surgical applications to animals or human patients? [ ] Yes , [ ] No ; if yes,  contact the UBC Radiation Safety Office prior to beginning any experiments.

- If applicable, provide detailed numbered steps taken to ensure persons are not exposed during laser demonstrations

- If applicable, provide detailed numbered steps taken to ensure persons are not exposed to laser hazards when fibre optics are used in association with the laser setup

- Itemize and describe control procedures for all non-laser hazards associated with this unit.

These may include but are not restricted to:

Electrical Hazards

- No one should work on lasers or power supplies unless qualified and approved to perform the specific tasks.

- Do not wear rings, watches, or other metallic apparel when working with electrical equipment.

- Do not handle electrical equipment when hands or feet are wet or when standing on a wet surface.

- When working with high voltages, regard all floors as conductive and grounded.

- Be familiar with electrocution rescue procedure and emergency first aid.

- Prior to working with electrical equipment, de-energize the power source. Lock and tag out the disconnect switch in accordance Lockout/Tagout.

- Check that each capacitor is discharged, shorted, and grounded prior to working in the area of the capacitors.

- When possible, use shock preventing shields, power supply enclosures, and  shielded leads in all experimental or temporary high voltage circuits.

Common Hazards Encountered When Working With Electrical Equipment

- Uncovered electrical terminals

- Improperly insulated electrical terminals

- Hidden power up/on warning lights

- Lack of personnel training in CPR (this and first aid training is offered at the Lab)

- Buddy system not being practiced during maintenance and alignment work

- Non earth-grounded/improperly grounded laser equipment

- Excessive wires and cables on the floor that create fall/trip hazards

Laser Dyes

Laser dyes are often toxic and/or carcinogenic chemicals dissolved in flammable solvents. This creates the potential for personnel exposures above permissible limits, fires, and chemical spills.

Frequently, the most hazardous aspect of a laser operation is the mixing of chemicals that make up the laser dye.  In addition, hazardous-waste-disposal concerns need to be addressed.

Compressed and Toxic Gases

Hazardous gases maybe used in laser applications, i.e., excimer lasers (fluorine, hydrogen chloride). The AHD should contain references for the safe handling of compressed gases, such as seminic restraints, use of gas cabinets, proper tubing and fittings, etc.

Cryogenic Fluids

Cryogenic fluids are used in cooling systems of certain lasers, and can create hazards situations. As these materials evaporate, they can create oxygen deficient atmospheres and an asphyxiation hazard by replacing the oxygen in the air. Adequate ventilation must be provided. Cryogenic fluids are potentially explosive when ice collects in valves or connectors that are not specifically designed for use with cryogenic fluids. Condensation of oxygen in liquid nitrogen presents a serious explosion hazard if the liquid oxygen comes in contact with any organic material. While the quantities of liquid nitrogen that may be used are usually small, protective clothing and face shields must be used to prevent freeze burns to the skin andeyes.

Radiofrequencies (RF)

Some lasers contain RF excited components as plasma tubes and Q switches. Unshielded and loosely tightened components may allow RF fields to leak from the device and expose staff.

Ergonomics

Ergonomic problems can arise from a laser operation by causing awkward unique arm and wrist positions.

If such repetitive deviations occur for prolong periods of time, medical problems such as repetitive strain injuries may arise

Seismic Safety

It is lab policy to prevent the loss of life and to minimize the risk of personal injury, program interruption, and property damage due to earthquakes. To support this goal, all laser users need to view their laser setup for compliance with seismic safety guidelines. Examples would be fastening electronic racks to the floor or walls, and racks on casters having at least two locking wheels. When possible, heavy laser equipment should be bolted down Fumes/Vapors/Laser Generated Air Contaminants (LGAC), from Beam / Target interaction

Air contaminants may be generated when certain Class 3b and Class 4 laser beams interact with matter.  When the target irradiance reaches a given threshold of approximately 10 to the 7th W/cm2 target materials, including plastics, composites, metals, and tissues, may liberate toxic and noxious airborne contaminants. In other words, when laser beams are sufficiently energized to heat up a target, the target

may vaporize, creating hazardous fumes or vapors that may need to be captured or exhausted.

When targets are heated to very high temperatures, as in laser welding and cutting, an intense bright light is emitted. This light often contains large amounts of short wavelength or blue light, which may cause conjunctivitis, photochemical damage to the retina, and/or erythema (sunburn-like reactions) in the skin.

Plasma Emissions

Interactions between very high power laser beams and target materials may in some cases produce plasmas. The plasma generated may contain hazardous “blue light” and UV emissions, which can be an eye and skin hazard.

UV and Visible Radiation

UV and visible radiation may be generated by laser discharge tubes and pump lamps. The levels

produced may be an eye and skin hazard.

Explosion Hazards

High pressure arc lamps, filament lamps, and capacitors may explode if they fail during operation. These

components are to be enclosed in a housing which will withstand the maximum explosive forces that may

be produced. Laser targets and some optical components also may shatter if heat cannot be dissipated

quickly enough. Consequently, care must be used to provide adequate mechanical shielding when

exposing brittle materials to high intensity lasers.

Ionizing Radiation (X-rays)

X-rays could be produced from two main sources, high voltage vacuum tubes of laser power supplies such as rectifiers, thyratrons, and electric discharge lasers. Any power supplies which require more than 15 kilovolts may produce enough x-rays to be a health concern.

LASER REGISTRATION DATA SHEET

Department____________________________________________________

Supervisor________________________________________  ph_________

Unit Location____________________________Rm___________________

Contact name_____________________________email_________________

Unit Description

Commercial Unit    Y  N

Manufacturer__________________________________

Model_________________________________________

Serial No.____________________________________

Class    III-b   IV

Wattage _____________    Wavelength Range___________________nm

Check:   Gas___ Solid State___  Semiconductor___  Liquid___

Continuous Wave  Y  N

or

Check:  Normal Pulse___ Long Pulse___

Q Pulsing___ (Mirror___ or Shutter___)

Enclosed Beam   Y  N  Partial